**Ultrasound Imaging Points to Line Registration.**

Alexis Boucharin, Csaba Pinter, Tamas Heffter, Andras Lasso

*Laboratory for Percutaneous Surgery, School of Computing, Queen’s University, Kingston, Ontario, Canada*

***Abstract***

Ultrasound imaging devices are now part of prostate brachytherapy preoperative planning and therefore have to be calibrated and tested as precisely as possible to help physicians elaborate their diagnosis and operative plan with accuracy. Image quality assurance and calibration of ultrasound imaging systems are performed on several types of phantoms containing fiducial lines in different configuration. Automatic fiducial line segmentation is therefore an important part of image quality assurance and calibration of ultrasound imaging systems. However, the current methods need to develop new code specific to a phantom. This operation is time consuming and the tolerance parameters are often fixed and obtained via a trial and error process. We propose a new method that will handle several configuration for the fiducial lines without requiring any change in the code, this method will detect coplanar lines that be contained in multiple planes for any number of lines per plane as well as any number of points per line. This method will also determine the tolerance parameters automatically from inputs such as maximum angular movement.

***Method***

The fiducial segmentation algorithm by (Chen, 2009) provides a list of fiducial points sorted by intensity that the method takes as an input.

From the list of fiducial points, N-points lines are computed and sorted by their intensity so that we have a list of lines each made of N fiducial points. The number or fiducial lines (fiducial points in a cross plane) is provided in the phantom configuration file, as well as their basic structures such as parallel fiducial lines and Z-shaped fiducial structure. The number of these structures is not limited and is provided in the phantom configuration file. Then, an exhaustive search will be performed on the N-points lines found previously to match the actual lines made from the fiducial points from the phantom configuration file. Once the lines are correctly matched, we can determine from image orientation and a transform matrix the correspondence between the fiducial points we found the actual one and therefore register them to the labels provided in the phantom configuration file.

Moreover, the different thresholds to accept points on a line or to register a potential line to an actual one will be computed by the algorithm instead of implemented directly after a trial and error process. From the angular maximum movements of the probe, we can determine how far from the actual position the candidate line can be. This angular maximum movement will provide the range in which the image can actually be as the image plane might not necessarily be perpendicular to the fiducial lines due to user movements, or could be slightly rotated to one side or the other. These angular parameters can be obtained from the phantom configuration file and the input data and would allow a less restrictive and more optimal choice of threshold parameters and will not need to be changed and found again for each phantom and for each use by the user.

The method has been developed in C++ with the **Insight Segmentation and Registration Toolkit** (ITK) and the Visualization Toolkit (VTK) for portability, speed and robustness. A Slicer3 module has been developed as well to provide a tool to visualize the results and the input data simultaneously in order to detect errors and the coherence of our results.

***Results and Discussion***

***References***

1. Chen, T. T. (2009). Chen, T.K., Thurston, A.D., Ellis, R.E., and Abolmaesumi, P. *Ultrasound in Med. & Biol, 35(1) pp. 79–93*.